## Slide 1: High speed PA testbench for studying vibrations on preloaded piezo actuators
Marta Alfonso
26-05-2023

## Slide 2: List of contents
Overview of the architecture of the system
Requirements of the system
Electrical design
Software design
  API
  GUI
Mechanical design
Measurements
System validation
Troubleshooting

## Slide 3: Architecture of the system
Functional generator
Power amplifier
Piezo stack
Power supply
PC

## Slide 4: Overview of the proposed solution
FGEN – 33600A Series Keysight Technologies
OPAMP - EK57U APEX
Piezo stack
PSU – SM70-22
Delta Elektronika
OPAMP - EK57U APEX
PSU – SM70-22 Delta Elektronika
PC

## Slide 5: Requirements of the system
Driving a piezoelectric brick up to 500 kHz with a full voltage range of 0 to 150 V.
User interface enabling remote generation of vibrations.
  API for configuration of the function generator that gives the input signal for the system.
  GUI for controlling and monitoring? the system from the PC.
A stand-alone chassis encapsulating the amplifier.
  Rack mountable
  Ventilation

## Slide 6: Electrical design: requirements
Voltage range [0, 150] V
Capacitive load (piezo brick) = 170 nF
Required slew rate (sine wave)
Required peak current

## Slide 7: Electrical design
Chosen amplifier: APEX MP111u

## Slide 8: Electrical design
Implemented with the APEX EK57 evaluation kit

## Slide 9: Electrical design
Unipolar bridge configuration with two amplifiers
  For a total voltage swing of 0 to 150 V
    Voltage swing of master amplifier
      0 to 75 V, Vpk = 37.5 V
    Voltage swing of slave amplifier
      0 to -75 V, Vpk = -37.5 V
  Required slew for each opamp (Vpk=37.5 V)
    117.8 V/µs
  Required peak current
    20 A

## Slide 10: Electrical design
Modified unipolar bridge configuration with two amplifiers
  Supply rails
    Report says +Vs = 85 V to have a maximum output of +Vs – 8.4 ≈ 75 V
    And –Vs =-20 V to have a minimum output of –Vs ≈ -14.2 V
    Input common mode range taking Vcm into account: +Vs – 15 V = 70 V and –Vs + 15 V = -5 V
    Datasheet:

## Slide 11: Electrical design: power supply
Modified unipolar bridge configuration with two amplifiers
  Supply rails
    +Vs = 85 V to have a maximum output of +Vs – 8.4 ≈ 75 V
    And –Vs =-20 V to have a minimum output of –Vs ≈ -14.2 V
  Available power supplies:
    2 x Delta Elektronika SM 70-22 → output range of 0-70 V, 22 A
    Delta Elektronika SM 35-45 → output range of 0-35 V, 45 A
    KEPCO BOP50 → output range of ±50 V, ± 4 A
65
20
+85
+20
65
20
-85
-20
20
20
+20
-20
70
70
+70
-70

## Slide 12: Electrical design: maximum values

## Slide 13: Software design
API for the remote configuration of the function generator.
  Python interface with Agilent 33600A series function generator.
  Use of PyVISA library with Keysight’s VISA driver for PC.
  Configuration of the instrument with SCPI commands.

## Slide 14: Software design
User interface
  Control of one or both input signals for the two amplifiers of the bridge configuration.
  Setting of vibration characteristics: Amplitude and Frequency.
  Visualization of the generated signals as well as signal across the piezo load.
  Logging and reporting features.

## Slide 15: Mechanical design
System on rack
  Encapsulation of the power amplifiers in a subrack
  Power supplies and generator mounted on the rack
Opamps subrack
  Front panel and back panel connections with other instruments
  ON/OFF general switch?
  Emergency stop button?
Preloaded piezo stacks mounting
  Piezo support? Mounting?
  Measurement system: interferometers?
Also needed to include in the rack:
  Measurement systems: vibrometer, interferometer?

## Slide 16: Mechanical design: rack configuration
+
+
620
483

## Slide 17: Mechanical design: devices
  We find three power supplies of the following model:

## Slide 18: Mechanical design: devices
  We find one power supply of the following model:
420,0
135,0
510,0

## Slide 19: Mechanical design: devices
  We find two power opamps that will need a subrack case and fans:
123,0
155,0
~110

## Slide 20: Mechanical design: devices
  We find one signal generator of the following model. Note that the grey protection is not present in the unit being used.
90
210
273

## Slide 21: Measurements
Measurements performed with the piezo stack mounted in the piezogoniometer?
Different/new structure for the testbench?
Moving part (piezo) vs fixed part (measurement system)
Relative or absolute measurements?
Which measurements?
  Angle of displacement? Distance of displacement?
  Speed? Acceleration?
  Frequency? Voltage amplitude?
Which instruments?
  Interferometer?

## Slide 22: System validation
Experiments to validate the system

## Slide 23: Troubleshooting
Problem #1: one of the amplifiers doesn’t work